Characteristics of the Terahertz Laser System

Typology of the system: photoconductive antennas for generation and detection of THz radiation, driven by a femtosecond laser system with 1560 nm center wavelength. Fully integrated electronics to control all components and to deliver THz spectroscopy measurements. Software to control THz system and image processing.

| Characteristic / parameter | Value |
| --- | --- |
| All components must be compatible with the Italian electrical system | YES |
| All components must be tested and perform according to terms in draft agreement | YES |
| CE marked components | YES |
| Laser power | ≥ 80 mW |
| Laser repetition rate | ≥ 100 MHz |
| Laser pulse duration | ≤ 60 fs |
| THz power generated by the antennas | ≥ 25 microWatts |
| Peak dynamic range THz signal in ≤ 1 minute data acquistion | ≥ 90 dB |
| Single-shot peak dynamic range (≤ 50 ms data acquisition) | ≥ 60 dB |
| THz bandwidth | ≥ 5 THz |
| Adjustable terahertz beam path length | ≥ 80 cm |
| Dynamic range at 5 THz | ≥ 20 dB |
| Electronics fully integrated (no external parts required) | YES |
| Imaging extension fully integrated | YES |
| Imaging software provided | YES |
| Translational movement of sample holder synchronized with THz data acquisition | YES |
| Accuracy of stage positioning | ≤ 0.3 mm |
| Measurement speed (with full spectrum recorded) | 16 pixels/s |
| Travel range of sample holder (in each in-plane direction) | ≥ 15 cm |
| Filtering of imaging data in time and frequency | YES |
| Delivery time | ≤ 3 months |